Supplementary File 1

Intravital Deep-Tumor Single-Beam 3-Photon, 4-Photon and Harmonic Microscopy

Gert-Jan Bakker1, Sarah Weischer1, Júlia Ferrer Ortas2, Judith Heidelin3, Volker Andresen3, Marcus Beutler4, Emmanuel Beaurepaire2 and Peter Friedl1,5,6

1 Department of Cell Biology, Radboud Institute for Molecular Life Sciences, Radboud University Medical Centre, 6525 GA Nijmegen, The Netherlands

2 Laboratory for Optics & Biosciences École Polytechnique, CNRS, INSERM, IP Paris, 91128 Palaiseau Cedex, France

3 LaVision BioTec GmbH, a Miltenyi Biotec company, 33617 Bielefeld, Germany

4 APE Angewandte Physik & Elektronik GmbH, 13053 Berlin, Germany

5 David H. Koch Center for Applied Genitourinary Cancers, The University of Texas MD Anderson Cancer Center, Houston, Texas 77030, USA

6 Cancer Genomics Centre, 3584 CG Utrecht, The Netherlands

Contact details corresponding authors:  Gert-Jan Bakker: email gert-jan.bakker@radboudumc.nl, P +31 (0)24 36 142 96. Peter Friedl: email peter.friedl@radboudumc.nl, P +31 (0)24 36 109 07. Mail address: Dept. of Cell Biology (283) RIMLS, Radboudumc, P.O. Box 9101, 6500 HB Nijmegen, The Netherlands.

Email addresses co-authors: Sarah Weischer: sarah.weischer@radboudumc.nl, Júlia Ferrer Ortas: julia.ferrer-ortas@polytechnique.edu, Judith Heidelin: heidelin@lavisionbiotec.de, Volker Andresen:  andresen@lavisionbiotec.de, Marcus Beutler: marcus_beutler@ape-berlin.de, Emmanuel Beaurepaire: emmanuel.beaurepaire@polytechnique.edu.

Supplementary File 1a. Comparison of parameters related to linear, lowIR and highIR excitation modalities.

| Modality | Confocal | lowIR (Ti:Sa/OPO) | 1300 nm highIR | 1650 nm highIR |
| --- | --- | --- | --- | --- |
| Processes | 1-photon, reflection | 2-, 3-photon, SHG, THG | 3-photon, SHG, THG | 3-, 4-photon, SHG, THG |
| Pulse frequency (MHz) | Continuous | 80 | 1 | 1 |
| Pulse length (fs) | - | 140 | 53 | 89 |
| Pulse energy focus (nJ) | - | < 0.4 | < 2 a | < 7 a |
| Peak power focus b | < 1 mW1,2 | < 3 kW | 38 kW | 78 kW |
| Power surface (mW) | < 12 | 40-120 | 2.8-33 | 8.7-38 |
| Emission I(E) … | E | E2, E3 | E2, E3 | E3, E4 |
| Pixel dwell time (µs) | 2 | 2-4 | 10-20 | 10-20 |
| Attenuation of light | very high | moderate | moderate | low |
| Water abs. sample | 0 % | < 3 % | < 6 % | < 26 % |
| Max. depth in tumor (µm) | < 1002,3 | 255 | 415 | 395 |

a)  Derived from maximum tolerable pulse energy (at the tumor surface, focus 50 µm deep in tissue, Figure 2d, h and i) and effective attenuation lengths for tumor tissue (Figure 3e).

b) Derived as pulse energy in the focus divided by pulse duration.

Supplementary File 1b. Order of the excitation processes (n) from emission intensity as a function of excitation energy (related to Figure 1).

| Exc. [nm] | Condition | Hoechst | eGFP | TagRFP | mCherry | AF680 | SHG | THG |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1650 | In vivo | 4.2±0.2 | -- | 3.2±0.2 | 3.1±0.4 | 3.0±0.3 | 2.15±0.04 | 3.2±0.1 |
| 1650 | Spheroids | -- | 3.8±0.3 | 3.0±0.2 | 3.0±0.1 | -- | -- | 3.0±0.2 |
| 1650 | Spheroidsa) | -- | -- | 3.1±0.1 | 3.0±0.2 | -- | -- | -- |
| 1300 | Spheroids | -- | 2.6±0.4 | 2.5±0.4 |  | -- | 1.85±0.08 | 2.9±0.3 |

a) Measured with the excitation source set to a repetition rate of 0.5 MHz.

Supplementary File 1c. Experimental parameters for brain measurements optimized for THG and/or AF680 emission.

| Channel / depth / Wavelength [nm] | Line av. | Port | DM | BP | PMT | Imm. | Seq. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| THG / < 830 µm / 1650 nm | 2 | 2 ch | 525/50 | no | GaAsP | H2O | 1 |
| THG / > 830 µm / 1650 nma) | 6 | 2 ch | 525/50 | no | GaAsP | H2O | 2 |
| THG / > 830 µm / 1650 nm | 6 | 2 ch | 880lp | no | GaAsP | D2O | 3 |
| AF680 / < 830 µm / 1650 nm | 2 | 4 ch | 900lp | 710/75 | GaAs | H2O | 1 |
| AF680 / > 830 µm / 1650 nm | 6 | 2 ch | 880lp | 710/75 | GaAs | H2O | 3 |
| THG / < 500 µm / 1270 nm | 3 | 2 ch | 485lp | 417/60 | GaAsP | H2O | 4 |
| AF680 / < 500 µm / 1270 nm | 3 | 4 ch | 900lp | 710/75 | Alkali | H2O | 4 |
| AF680 / 500-669 µm / 1270 nm | 1 | 4 ch | 900lp | 710/75 | Alkali | H2O | 5 |
| AF680 / > 669 µm / 1270 nm | 3 | 4 ch | 900lp | 710/75 | Alkali | H2O | 6 |

a) Solely used for generation of the 1650 nm signal attenuation curve.

Additional References

1.	Pawley, J. B. Handbook Of Biological Confocal Microscopy. (Springer, 2006).

2.	Centonze, V. E. & White, J. G. Multiphoton excitation provides optical sections from deeper within scattering specimens than confocal imaging. Biophys. J. 75, 2015–2024 (1998).

3.	Helmchen, F. & Denk, W. Deep tissue two-photon microscopy. Nat. Methods 2, 932–940 (2005).