## rendements en profondeur clinac
Figure 1-Figure supplement 1-source data 1
Rendement en profondeur du Clinac du Centre Oscar Lambret
Maxime Tomezak et Thomas Lacornerie
BEGIN_SCAN_DATA
| FORMAT=CC-Export | V1,9
| FILE_CREATION_DATE=11-May-2015 | 19:26
| LAST_MODIFIED=11-May-2015 | 19:26
| BEGIN_SCAN | 1
| TASK_NAME=tba | PDD | Profiles
PROGRAM=tbaScan
| MEAS_DATE=11-May-2015 | 19:24
| LINAC=CLINAC | 2
MODALITY=X
ISOCENTER=1000,00
INPLANE_AXIS=Y
CROSSPLANE_AXIS=X
DEPTH_AXIS=Depth
INPLANE_AXIS_DIR=GUN_TARGET
CROSSPLANE_AXIS_DIR=LEFT_RIGHT
DEPTH_AXIS_DIR=UP_DOWN
ENERGY=20,00
NOMINAL_DMAX=32,30
SSD=1000,00
SCD=450,00
BLOCK=0
WEDGE_ANGLE=0,00
FIELD_INPLANE=100,00
FIELD_CROSSPLANE=100,00
FIELD_TYPE=RECTANGULAR
GANTRY=0,00
GANTRY_UPRIGHT_POSITION=0
GANTRY_ROTATION=CW
COLL_ANGLE=0,00
COLL_OFFSET_INPLANE=0,00
COLL_OFFSET_CROSSPLANE=0,00
SCAN_DEVICE=MP3
SCAN_DEVICE_SETUP=BARA_LEFT_RIGHT
ELECTROMETER=TANDEM
RANGE_FIELD=AUTO
RANGE_REFERENCE=AUTO
DETECTOR=THIMBLE_CHAMBER
DETECTOR_SUBCODE=SEMIFLEX
DETECTOR_RADIUS=2,75
| DETECTOR_NAME=PTW | 31010 | Semiflex
DETECTOR_SN=801
DETECTOR_CALIBRATION=293100000,00
DETECTOR_IS_CALIBRATED=1
DETECTOR_REFERENCE=THIMBLE_CHAMBER
DETECTOR_REFERENCE_SUBCODE=SEMIFLEX
DETECTOR_REFERENCE_RADIUS=2,75
| DETECTOR_REFERENCE_NAME=PTW | 31013 | Semiflex
DETECTOR_REFERENCE_SN=1443
DETECTOR_REFERENCE_IS_CALIBRATED=1
DETECTOR_REFERENCE_CALIBRATION=93960000,00
REF_FIELD_DEPTH=0,00
REF_FIELD_DEFINED=ISOCENTER
REF_FIELD_INPLANE=100,00
REF_FIELD_CROSSPLANE=100,00
REF_SCAN_POSITIONS=0,00;1,50;3,00;4,50;6,00;7,50;9,00;10,50;12,00;13,50;15,00;16,50;18,00;19,50;21,00;22,50;24,00;25,50;27,00;28,50;30,00;31,50;33,00;34,50;35,00;37,50;40,00;42,50;45,00;47,50;50,00;52,50;55,00;57,50;60,00;62,50;65,00;67,50;70,00;72,50;75,00;80,00;85,00;90,00;95,00;100,00;105,00;110,00;115,00;120,00;125,00;130,00;135,00;140,00;145,00;150,00;155,00;160,00;165,00;170,00;175,00;180,00;185,00;190,00;195,00;200,00;205,00;210,00;215,00;220,00;225,00;230,00;235,00;240,00;245,00;250,00;255,00;260,00;265,00;270,00;275,00;280,00;285,00;290,00;295,00;300,00;
SCAN_CURVETYPE=PDD
SCAN_OFFAXIS_INPLANE=0,00
SCAN_OFFAXIS_CROSSPLANE=0,00
SCAN_ANGLE=0,00
SCAN_DIAGONAL=NOT_DIAGONAL
SCAN_DIRECTION=NEGATIVE
MEAS_MEDIUM=WATER
MEAS_PRESET=REFERENCE_DOSEMETER
MEAS_TIME=0,200
MEAS_UNIT=A,U,
| SCAN_SPEEDS=20,00; | 8,00;40,00; | 25,00;100,00; | 40,00;400,00; | 50,00;
| DELAY_TIMES=20,00; | 0,000;150,00; | 0,000;400,00; | 0,000;
PRESSURE=1013,20
TEMPERATURE=20,00
NORM_TEMPERATURE=20,00
CORRECTION_FACTOR=1,0000
EXPECTED_MAX_DOSE_RATE=3,00
BEGIN_DATA
| 0 | 1.4119 | 1.8367 | 34.7759 | 56.1682 | 0 | 34.7759 | 56.1682
| 1.5 | 1.6083 | 1.995 | 39.6133 | 61.0092 | 1.5 | 39.6133 | 61.0092
| 3 | 2.0843 | 2.2804 | 51.3374 | 69.737 | 3 | 51.3374 | 69.737
| 4.5 | 2.5023 | 2.5405 | 61.633 | 77.6911 | 4.5 | 61.633 | 77.6911
| 6 | 2.8017 | 2.7457 | 69.0074 | 83.9664 | 6 | 69.0074 | 83.9664
| 7.5 | 3.0087 | 2.9092 | 74.1059 | 88.9664 | 7.5 | 74.1059 | 88.9664
| 9 | 3.141 | 3.012 | 77.3645 | 92.1101 | 9 | 77.3645 | 92.1101
| 10.5 | 3.29 | 3.0931 | 81.0345 | 94.5902 | 10.5 | 81.0345 | 94.5902
| 12 | 3.4179 | 3.157 | 84.1847 | 96.5443 | 12 | 84.1847 | 96.5443
| 13.5 | 3.5301 | 3.2002 | 86.9483 | 97.8654 | 13.5 | 86.9483 | 97.8654
| 15 | 3.6457 | 3.2245 | 89.7956 | 98.6086 | 15 | 89.7956 | 98.6086
| 16.5 | 3.7471 | 3.2398 | 92.2931 | 99.0765 | 16.5 | 92.2931 | 99.0765
| 18 | 3.8278 | 3.2624 | 94.2808 | 99.7676 | 18 | 94.2808 | 99.7676
| 19.5 | 3.8806 | 3.2698 | 95.5813 | 99.9939 | 19.5 | 95.5813 | 99.9939
| 21 | 3.9298 | 3.2676 | 96.7931 | 99.9266 | 21 | 96.7931 | 99.9266
| 22.5 | 3.9675 | 3.2658 | 97.7217 | 99.8716 | 22.5 | 97.7217 | 99.8716
| 24 | 3.9973 | 3.2586 | 98.4557 | 99.6514 | 24 | 98.4557 | 99.6514
| 25.5 | 4.0155 | 3.2549 | 98.9039 | 99.5382 | 25.5 | 98.9039 | 99.5382
| 27 | 4.0406 | 3.2338 | 99.5222 | 98.893 | 27 | 99.5222 | 98.893
| 28.5 | 4.0514 | 3.2322 | 99.7882 | 98.844 | 28.5 | 99.7882 | 98.844
| 30 | 4.0543 | 3.2153 | 99.8596 | 98.3272 | 30 | 99.8596 | 98.3272
| 31.5 | 4.0608 | 3.1889 | 100.0197 | 97.5199 | 31.5 | 100.0197 | 97.5199
| 33 | 4.0617 | 3.1513 | 100.0419 | 96.37 | 33 | 100.0419 | 96.37
| 34.5 | 4.0649 | 3.1213 | 100.1207 | 95.4526 | 34.5 | 100.1207 | 95.4526
| 35 | 4.0512 | 3.0837 | 99.7833 | 94.3028 | 35 | 99.7833 | 94.3028
| 37.5 | 4.0488 | 3.0536 | 99.7241 | 93.3823 | 37.5 | 99.7241 | 93.3823
| 40 | 4.0348 | 3.0161 | 99.3793 | 92.2355 | 40 | 99.3793 | 92.2355
| 42.5 | 4.0155 | 2.9822 | 98.9039 | 91.1988 | 42.5 | 98.9039 | 91.1988
| 45 | 3.9975 | 2.9372 | 98.4606 | 89.8226 | 45 | 98.4606 | 89.8226
| 47.5 | 3.9654 | 2.9127 | 97.67 | 89.0734 | 47.5 | 97.67 | 89.0734
| 50 | 3.9346 | 2.8756 | 96.9113 | 87.9388 | 50 | 96.9113 | 87.9388
| 52.5 | 3.9053 | 2.8413 | 96.1897 | 86.8899 | 52.5 | 96.1897 | 86.8899
| 55 | 3.8855 | 2.8129 | 95.702 | 86.0214 | 55 | 95.702 | 86.0214
| 57.5 | 3.8489 | 2.7793 | 94.8005 | 84.9939 | 57.5 | 94.8005 | 84.9939
| 60 | 3.8115 | 2.7398 | 93.8793 | 83.7859 | 60 | 93.8793 | 83.7859
| 62.5 | 3.7829 | 2.7039 | 93.1749 | 82.6881 | 62.5 | 93.1749 | 82.6881
| 65 | 3.7544 | 2.6681 | 92.4729 | 81.5933 | 65 | 92.4729 | 81.5933
| 67.5 | 3.7183 | 2.6061 | 91.5837 | 79.6972 | 67.5 | 91.5837 | 79.6972
| 70 | 3.6818 | 2.5369 | 90.6847 | 77.581 | 70 | 90.6847 | 77.581
| 72.5 | 3.6445 | 2.4734 | 89.766 | 75.6391 | 72.5 | 89.766 | 75.6391
| 75 | 3.6121 | 2.4087 | 88.968 | 73.6606 | 75 | 88.968 | 73.6606
| 80 | 3.5493 | 2.3504 | 87.4212 | 71.8777 | 80 | 87.4212 | 71.8777
| 85 | 3.4727 | 2.2883 | 85.5345 | 69.9786 | 85 | 85.5345 | 69.9786
| 90 | 3.4141 | 2.2257 | 84.0911 | 68.0642 | 90 | 84.0911 | 68.0642
| 95 | 3.3429 | 2.161 | 82.3374 | 66.0856 | 95 | 82.3374 | 66.0856
| 100 | 3.2827 | 2.109 | 80.8547 | 64.4954 | 100 | 80.8547 | 64.4954
| 105 | 3.2132 | 2.0477 | 79.1429 | 62.6208 | 105 | 79.1429 | 62.6208
| 110 | 3.1505 | 2 | 77.5985 | 61.1621 | 110 | 77.5985 | 61.1621
| 115 | 3.088 | 1.9438 | 76.0591 | 59.4434 | 115 | 76.0591 | 59.4434
| 120 | 3.0328 | 1.8923 | 74.6995 | 57.8685 | 120 | 74.6995 | 57.8685
| 125 | 2.9682 | 1.8427 | 73.1084 | 56.3517 | 125 | 73.1084 | 56.3517
| 130 | 2.9129 | 1.7901 | 71.7463 | 54.7431 | 130 | 71.7463 | 54.7431
| 135 | 2.8537 | 1.7447 | 70.2882 | 53.3547 | 135 | 70.2882 | 53.3547
| 140 | 2.799 | 1.6949 | 68.9409 | 51.8318 | 140 | 68.9409 | 51.8318
| 145 | 2.7448 | 1.6472 | 67.6059 | 50.3731 | 145 | 67.6059 | 50.3731
| 150 | 2.6851 | 1.6032 | 66.1355 | 49.0275 | 150 | 66.1355 | 49.0275
| 155 | 2.6332 | 1.5625 | 64.8571 | 47.7829 | 155 | 64.8571 | 47.7829
| 160 | 2.5783 | 1.5161 | 63.5049 | 46.3639 | 160 | 63.5049 | 46.3639
| 165 | 2.5294 | 1.4757 | 62.3005 | 45.1284 | 165 | 62.3005 | 45.1284
| 170 | 2.4853 | 1.434 | 61.2143 | 43.8532 | 170 | 61.2143 | 43.8532
| 175 | 2.4342 | 1.3968 | 59.9557 | 42.7156 | 175 | 59.9557 | 42.7156
| 180 | 2.3835 | 1.3534 | 58.7069 | 41.3884 | 180 | 58.7069 | 41.3884
| 185 | 2.3375 | 1.3159 | 57.5739 | 40.2416 | 185 | 57.5739 | 40.2416
| 190 | 2.2917 | 1.2814 | 56.4458 | 39.1865 | 190 | 56.4458 | 39.1865
| 195 | 2.2438 | 1.2444 | 55.266 | 38.055 | 195 | 55.266 | 38.055
| 200 | 2.2033 | 1.2103 | 54.2685 | 37.0122 | 200 | 54.2685 | 37.0122
| 205 | 2.1582 | 1.1783 | 53.1576 | 36.0336 | 205 | 53.1576 | 36.0336
| 210 | 2.1126 | 1.1447 | 52.0345 | 35.0061 | 210 | 52.0345 | 35.0061
| 215 | 2.0723 | 1.1123 | 51.0419 | 34.0153 | 215 | 51.0419 | 34.0153
| 220 | 2.0328 | 1.083 | 50.069 | 33.1193 | 220 | 50.069 | 33.1193
| 225 | 1.9899 | 1.0546 | 49.0123 | 32.2508 | 225 | 49.0123 | 32.2508
| 230 | 1.9558 | 1.0209 | 48.1724 | 31.2202 | 230 | 48.1724 | 31.2202
| 235 | 1.9157 | 0.9923 | 47.1847 | 30.3456 | 235 | 47.1847 | 30.3456
| 240 | 1.8793 | 0.96709 | 46.2882 | 29.5746 | 240 | 46.2882 | 29.5746
| 245 | 1.8434 | 0.93883 | 45.4039 | 28.7104 | 245 | 45.4039 | 28.7104
| 250 | 1.8074 | 0.9182 | 44.5172 | 28.0795 | 250 | 44.5172 | 28.0795
| 255 | 1.7737 | 0.89024 | 43.6872 | 27.2245 | 255 | 43.6872 | 27.2245
| 260 | 1.7347 | 0.86609 | 42.7266 | 26.4859 | 260 | 42.7266 | 26.4859
| 265 | 1.7007 | 0.84456 | 41.8892 | 25.8275 | 265 | 41.8892 | 25.8275
| 270 | 1.6686 | 0.82072 | 41.0985 | 25.0985 | 270 | 41.0985 | 25.0985
| 275 | 1.6362 | 0.79708 | 40.3005 | 24.3755 | 275 | 40.3005 | 24.3755
| 280 | 1.6056 | 0.77779 | 39.5468 | 23.7856 | 280 | 39.5468 | 23.7856
| 285 | 1.5761 | 0.75926 | 38.8202 | 23.219 | 285 | 38.8202 | 23.219
| 290 | 1.5443 | 0.73694 | 38.0369 | 22.5364 | 290 | 38.0369 | 22.5364
| 295 | 1.5141 | 0.71515 | 37.2931 | 21.87 | 295 | 37.2931 | 21.87
| 300 | 1.4877 | 0.738 | 36.6429 | 22.5688 | 300 | 36.6429 | 22.5688
END_DATA
| END_SCAN | 1
END_SCAN_DATA