| Problem | Cause | Solution |
| --- | --- | --- |
| - Burning of electronic components (ZAF specific). - Food distribution is non homogeneous (ZAF specific). - All the pumps connected to a motor controller (LN298) are not working. - A lot of food remains in the food preparation tank. - No food going inside the food preparation tank. - Water leak at the pump output. - Dirty fish water feeding . - Tubes are getting dirty quickly. - The mixed water/food is not going out of the tube - Raspberry Pi is not responsive | - Voltage ripples due. - tubing issues. - Wiring problem. - bad pumping. - Food delivery is clogged. - Tubing connections with the pump. - Overfeeding. - Cleaning is not well done. - The pump is not running correctly and our tubes are too long. | - Add 1000uF capacitors in parallel to the motor terminals - All the tubes should be same length - One tube might be pinched or bent. - Check the T splice connector wiring connections - Check the tube position inside the food preparation container, it must be at the bottom. - Possible pump malfunction, change the pump. - Clean the funnel and the food dispenser output. - Change the zip tie and eventually cut the tip of the damaged tube. - Reduce the food quantity delivered by i. Adjusting the software, ii. Adjust the food dispenser closure. - Add more cleaning programs. - Possible pump malfunction, change the pump. - Increase the pump running period in the code source. |